Use

Kit for measurement of uric acid in serum, plasma and urine

Colorimetric enzymatic method Uricase-POD-PAP.

Summary

Uric acid measurements are used in the diagnosis and treatment of gout and impaired renal fuction.

Principle

End point analysis. Uric acid is converted by uricase and H2O2  which, under the catalytic influence of peroxide (POD), oxidizes compound, reacts with 4-aminophenazone and 3,5-diclorophenol-sulphonate giving a red coloured compound. The increase in absorbance generated by the red dye is proportional to the uric acid concentration in the sample.

Reagents

R1 	Goods buffer   pH 8.0          100.0 mmol/l

ascorbate oxidase		³ 200 U/l

3,5-diclorophenol-sulphonate  2.5 mmol/l

Goods buffer  pH 8.0           100.0 mmol/l

4-aminophenazone                 0.3 mmol/l

Peroxidise	                ³ 3000 U/l

Uricase	                ³   600 U/l

Reagents Preparation

Reagents are liquid and ready to use.

Storage And Stability

- 	Store the kit at 2-8°C.

- 	After opening, the vial R1 is stable 90  days if recapped immediately and protected from contamination, evaporation, direct light, and stored at the correct temperature.

.

Precaution in Use

The product is not classified as dangerous (DLg. N. 285 art. 28 l. n. 128/1998).

However the reagent should be handled with care, according to good laboratory practice.

Caution: the reagents contain Sodium Azide (0.095%) as preservative. Avoid swallowing and contact with skin, eyes and mucous membranes.

Waste Management

Please refer to the local legal requirements.

Specimen Collection and Preparation

- 	Serum or heparinized plasma.

-  	Diluted urine 1:10

- 	Do not use samples with haemolysis.

- 	Do not use oxalate as anticoagulant. It’s not advisable using EDTA and fluoride because they could cause positive interferences.

- 	The uric acid in the serum and in the plasma is stable for 3 days if kept at 2-8°C or 6 months at -20°C.

Note

- 	The kit, according to this method, must be used in manual procedures. About automatic use follow specific applications.

- 	Avoid direct light, contamination and evaporation.

- 	The volumes in the procedure can be changed proportionally.

- 	In case of complaint or quality control request, refer to the lot number on the package or the lot number on the singles vials.

Procedures

Wavelength	 λ: 550(520-550) nm

Working temperature 	37°C

Optical path  	1 cm

Reaction 	“end point” (increasing)

Bring the reagents at 15-25°C before use them.

Monoreagent Procedure “sample starter”

BLANK         STD     SAMPLE

Reagent 	1000 µl 	   1000 µl 	1000 µl

Distilled Water 	25 µl 	-- 	--

Sample 	-- 	-- 	25 µl

Standard 	-- 	25 µl 	--

Mix, then incubate for 5’ at 37 °C. Measure the absorbance of sample (EC) and standard (ESTD) against the reagent blank.

Calculation

Uric acid (mg/dl) or (µmol/l) =

EC/ESTD x Conc. STD

Diluted urines: multiply the result for diluting factor.

Conversion Factor

Uric acid [mg/dl] x  59.48  =  Uric acid [µmol/l]

Reference Values to 37°C

Serum - plasma

Women	  2.4 –	 5.7 	mg/dl  	( 143 – 339 µmol/l )

Men       	3.4 –	7.0 	mg/dl 	( 202 – 416 µmol/l )

Urine:   250 – 	750 mg/24h 	(1.49 – 4.46 mmol/24h)

Reference values are considered indicative since each laboratory should establish reference ranges for its own patient population. The analytical results should be evaluated  with other information coming from patient’s clinical history.

ANALYTICAL PERFORMANCES

Linearity

Reaction is linear up to a concentration of 25 mg/dl. Samples with values exceeding this range must be diluted with saline solution. Then, multiply the result for diluting factor.

Analytical Sensitivity

The test sensitivity in terms of detection limit is 0.50 mg/dl.

“Intra-Assay” precision (within-Run)

Determined on 20 samples for each control (N-H) (Normal-High). Results:

MEAN (mg/dl)	N = 4.40	 H = 12.64

S.D.		N = 0.14 	 H =   7.66

C.V.%		N = 3.09  	 H =   2.61

“Inter-Assay” precision (between-run)

Determined  on 20 samples for each control (N-H). Results:

MEAN (mg/dl)	N = 4.48	H = 10.89

S.D.		N = 0.12	H =   0.24

C.V.%		N = 2.77	H =   2.23

Correlation

A study based comparing this method with a similar method on 20 samples has given a correlating factor r = 0.99

y = 1.00x + 0.2737

Interferences

No interferences was observed by the presence of :

Bilirubin 	≤ 15 	mg/dl

Triglycerides 	≤ 800	mg/dl

Ascorbate acid 	≤ 20 	mg/dl

Hemoglobin	≤ 50 	mg/dl

Lipemic specimens should not be used for analysis.For a comprehensive review of interfering substances, refer to the publication by Young.

Quality Controls

It’s necessary, each time the kit is used, to perform the quality controls and to check that values obtained are within the acceptance

range provided in the insert. Each laboratory should estabilish its own mean and standard deviation and adopt a quality control program to monitor laboratory testing.

Bibliography

Kaplan LA, Pesce AJ: “Clinical Chemistry”, Mosby Ed. (1996). Barham D, Trinder P: Analyst, 97 142 (1972).

Fossati P, Prencipe L, Berti G: Clin. Chem., 26(2) 227 (1980).

Young D.S., Effects of Drugs on Clinical Laboratory Tests,AACC Press, Washington, DC 5th ed.2000.

Symbols

|  | CE Mark (98/79 CE regulation) |
| --- | --- |
|  | in vitro medical device |
|  | Batch Code |
|  | Use by |
|  | Storage temperature limits |
|  | Read instruction for use |
|  | Gesan Production srl |