Use

Kit for measurement of total bilirubin in serum or plasma.Colorimetric  method Jendrassik -Grof modified.

Summary

Total bilirubin measurements are used in the diagnosis and treatments of  various liver diseases, and  haemolytic and haematology disorders.

Principle

End point analysis. Total bilirubin reacts with sulphanilic acid, giving a coloured azocompound (azobilirubin). The increase in absorbance due to the formation of azobilirubin is proportional to the total bilirubin concentration in the sample.

Reagents

R1	Sulphanilic acid    12.5   mmol/l

citric acid    	          1.00     mol/l

caffeine                0.18      mol/l

preservatives and surface-active agents not anionic

R2	Nitrite sodium      0.52  mmol/l

Reagent Preparation

Reagents are liquid and ready to use. About using as monoreagent (“sample-starter”) add to every 4 ml of R1 reagent, 1 ml of R2 reagent.

Storage  And Stability

- 	Store the kit at 15-25°C.

- 	After opening, the vials R1 and R2 are stable 90 days if recapped immediately and protected from contamination, evaporation, direct light, and stored at the correct temperature.

- 	Working solution stability (R1+ R2): 14 days at 2-8°C.

Precaution In Use

The product is not classified as dangerous (DLg. N. 285 art. 28 l. n. 128/1998). The final concentration of the components, is below the limits imposed by Directives 67/548/EEC and 88/379/EEC and subsequent amendments-to the classification packaging and labeling of dangerous substances.

However the reagent should be handled with caution, according to good laboratory practice.

Caution: the reagents contain Caffeine. Avoid swallowing and contact with skin, eyes and mucous membranes.

Waste Management

Please refer to the local legal requirements.

Sample

-  Serum or EDTA-Na2 plasma.

-  Do not use samples with haemolysis.

-  Do not use lipemic sample.

- Keep the samples far from light and heat        because the bilirubin is a photosensitive pigment.

- Perform the samples as soon as possible.

Note

- 	The kit, according to this method, must be used in manual procedures. About automatic using follow specific applications.

- 	Avoid direct light, contamination and evaporation.

- 	The volumes in the procedure can be changed proportionally.

- 	In case of complaint or quality control request, refer to the lot number on the package or the lot  number on the singles vials.

Procedure

| Wavelength | λ: 570 (550 – 580) nm |
| --- | --- |
| Working Temperature | 37°C |
| Optical Path | 1 cm |
| Reaction | “end point” |

Monoreagent Procedure “sample starter”

|  | Blank |  | Sample |
| --- | --- | --- | --- |
| Working Reagent | 1500μl |  | 1500μl |
| Distilled Water | 100 μl |  | – |
| Sample | – |  | 100 μl |
| Mix, then incubate 5’ at 37°C. Measure the absorbance of sample (EC) and blank (EBC) against water. |  |  |  |

Bireagent Procedure “substrate starter”

|  | Blank |  | Sample |
| --- | --- | --- | --- |
| Reagent R1 | 1200μl |  | 1200μl |
| Distilled Water | 100 μl |  | – |
| Sample | – |  | 100 μl |
| Mix, incubate at 37°C for 5’ and then add: Reagent R2 300 μl 300 μl |  |  |  |
| Mix, then incubate 5’ at 37°C. Measure the absorbance of sample (EC) and blank (EBC) against water. |  |  |  |

Calculation

Total bilirubin  [ mg/dl ] = (EC-ECB) x 14.5

Total bilirubin  [ μmol/l] = (EC-ECB) x  248

The reagents performances are related to 37°C, 1 cm and 570 nm.

Conversion Factor

Bilirubin [mg/dl]  x  17.0  =  Bilirubin [µmol/l]

Reference Values

Adults       0.3 – 1.2  mg/dl (5.1 -  17.1 μmol/l)

Reference values are considered indicative since each laboratory should establish reference ranges for its own patient population. The analytical results should be evaluated  with other information coming from patient’s clinical history.

Analytical Performances

The performance of the reagent are related to 37 ° C, 1 cm and 570 nm

Linearity

The reaction is linear in concentration range between 0.04 e 25 mg/dl. Samples with values exceeding 25 mg/dl must be diluted with saline solution. Then, multiply, , the result for diluting factor.

“Intra-Assay” precision (within-Run)

Determined on 20 samples for each control (N-H) (Normal-High). Results:

| MEAN (mg/dl) | N = 1.32 | H = 4.10 |
| --- | --- | --- |
| S.D. | N = 0.03 | H = 0.11 |
| C.V.% | N = 2.38 | H = 2.71 |

“Inter-Assay” precision (between-Run)

Determined on 20 samples for each control      (N-H).Results:

| MEAN (mg/dl) | N = 1.32 | H = 4.01 |
| --- | --- | --- |
| S.D. | N = 0.04 | H = 0.08 |
| C.V.% | N = 2.71 | H = 1.89 |

Analytical sensitivity

The test sensitivity in terms of detection limit is 0.04 mg/dl.

Correlation

A study based comparing this method with a similar method on 2 samples has given a correlating factor r = 0.99

y = 1.0077x + 0.034

Interferences

No interference was observed by the presence of

Triglycerides	≤ 200 mg/dl

Haemoglobin	≤ 500 mg/dl

For a comprehensive review of interfering substances, refer to the publication by Young.

Quality Controls

It’s necessary, every time the kit is used, to make the quality controls and to check that values obtained are within the acceptance range provided in the insert. Each laboratory should establish its own mean and standard deviation and adopt  a quality control program to monitor laboratory testing.

Bibliography

Kaplan, L.A., Pesce, A..J.: “Clinical Chemistry”, Mosby Ed. (1996).

Jendrassik, L., Gróf, P., Biochem. Z., 297, 81 (1938).

Jakobs, D.S., Kasten, Jr. B.L., Demmott, W.R., Wolfson, W.L.: “Laboratory Test Handbook”, Lexi-Comp and Williams & Wilkins Ed. 2nd Edition (1990).

Young D.S., Effects of Drugs on Clinical Laboratory Tests, AACC Press, Washington, DC 5th ed.2000.

Simbols

|  | CE Mark (requirement of 98/79 regulation) |
| --- | --- |
|  | in vitro medical device |
|  | Batch Code |
|  | Use by |
|  | Storage temperature limits |
|  | Read instruction for use |
|  | Gesan Production srl |