Use

Kit for measurement of creatinkinase in serum or plasma.

Kinetic optimized method DGKC-IFCC.

Summary

CK measurements are used in the diagnosis and treatment of  myocardial infarction and muscle diseases such as progressive, Duchenne – type muscular dystrophy.

Principle

Kinetic analysis. Creatinkinase catalyzes reaction between creatine phosphate and ADP, giving creatine and ATP. This one, in presence of glucose and hexokinase is transformed into ADP and glucose-6-phosphate which, with glucose-6-phosphate-dehyidrogenase intervention, forms glucose-6-phosphogluconate while  NADP+ is reduced to NADPH. The CK activity in the sample is calculated measuring the absorbance variation, due to the transformation of NADP+ into NADPH.

Reagents

R1	Imidazole buffer	pH 6.3 150.0 mmol/l

AMP   	25.0 mmol/l

NADP   	2.0 mmol/l

D-glucose   	25.0 mmol/l

diadenosine phosphate	12.0 mmol/l

hexokinase    	3000 U/l

magnesium acetate   	12.0 mmol/l

N-acetyl-cisteine 	25.0 mmol/l

R2	Imidazole buffer  	pH 6.3  150.0 mmol/l

ADP  	12.0 mmol/l

G6P-DH   	10000 U/l

creatine phosphate   	35.0 mmol/l

Reagents Preparation

Reagents are liquid and ready to use. About using as monoreagent (“sample-starter” procedure)  add to every 4 ml of R1 reagent, 1 ml of R2 reagent. Keep out the reagents from refrigerator only for the use and recap them immediately.

Storage And Stability

- 	Store the kit at 2-8°C.

- 	After opening, the vials R1 and R2 are stable 90 days if recapped immediately and protected from contamination, evaporation, direct light, and stored at the correct temperature.

- 	Working solution stability (R1+ R2): 20 days at 2-8°C.

Precaution in Use

The product is not classified as dangerous (DLg. N. 285 art. 28 l. n. 128/1998). The final concentration of the components, is below the limits imposed by Directives 67/548/EEC and 88/379/EEC and subsequent amendments-to the classification packaging and labeling of dangerous substances.

However the reagent should be handled with care, according to good laboratory practice.

Caution: the reagents contain Sodium Azide (0.095%) as preservative. Avoid swallowing and contacting with skin, eyes and mucous membranes.

Waste Management

Please refer to the local legal requirements.

Specimen Collection and Preparation

-   Serum or heparinized plasma.

-  Store the samples protected from light.

-  The CK activity is stable for 7 days at 2-8°C or 30 days at -20°C

Note

- 	The kit, according to this method, must be used in manual procedures. About automatic use follow specific applications.

- Avoid direct light, contamination and evaporation.

- 	The volumes in the procedure can be changed proportionally.

- 	In case of complaint or quality control request, refer to the lot number on the package or the lot number on the singles vials.

Procedures

Wavelength	 λ: 340 (330 o 370) nm

Working temperature 	37°C

Optical path	 1 cm

Reaction 	kinetic (increasing)

Bring the reagents at 15-25°C before using them.

Monoreagent Procedure  “sample starter”

|  | BLANK | SAMPLE |
| --- | --- | --- |
| Working Reagent | 1000 µl | 1000 µl |
| Distilled Water | 25 µl | -- |
| Sample | -- | 25 µl |
| Mix, then incubate for 2’ a 37°C. Measure the absorbance of sample (EC). Make at least two readings at a distance of 60’’. Calculate the absorbance variation ΔE/min from performed readings. |  |  |
| Bireagent Procedure “substrate starter” |  |  |
|  | BLANK | SAMPLE |
| Reagent R1 | 800 µl | 800 µl |
| Distilled Water | 25 µl | -- |
| Sample | -- | 25 µl |
| Mix, then incubate for 2’ a 37°C. Then add: Reagent R2 200 µl 200 µl Mix , then incubate for 2’ a 37°C. Measure the absorbance of sample (EC). Make at least two readings at a distance of 60’’. Calculate the absorbance variation ΔE/min from performed readings. |  |  |
| Calculation CK NAC [U/l] = ΔE/min x 6507 The factor and the reagent performances are related to 37°C, 1 cm and 340 nm. |  |  |

Reference Values

| Siero | Plasma (U/l) | 37°C |
| --- | --- | --- |
| Adults | Women | 24 – 170 U/l |
|  | Men | 24 – 195 U/l |
| Children | Babies | 468 – 1200 U/l |
|  | ≤ 5 days | 195 – 700 U/l |
|  | < 6 months | 41 – 330 U/l |
|  | > 6 months | 24 – 229 U/l |

Reference values are considered indicative since each laboratory should establish reference ranges for its own patient population. The analytical results should be evaluated  with other information coming from patient’s clinical history.

ANALYTICAL PERFORMANCES

The performance of the reagent are related to 37 ° C, 1 cm and 340 nm

Linearity

The reaction is linear in concentration range between 10 e 1000 U/l. Samples with values exceeding 1000 U/l  must be diluted with saline solution. Then multiply the result for diluting factor.

Analytical Sensitivity

The test sensitivity in terms of detection limit is 10 U/l.

“Intra-Assay” precision (within-Run)

Determined on 20 samples for each control (N-H) (Normal-High). Results:

MEAN (U/l)	N = 150.75	 H = 450.85

S.D.	N =     2.21	 H =     3.76

C.V.%	N =     1.47	 H =     0.83

“Inter-Assay” precision (between-run)

Determined  on 20 samples for each control (N-H). Results:

MEAN  (U/l)	N = 152.38	H = 449.18

S.D.	N =     2.77	H =     5.42

C.V.%	N =     1.82	H =     1.21

Correlation

A study based comparing this method with a similar method on 20 samples has given a correlating factor r = 0.99

y = 1.0181 x + 5.247

Interferences

No interferences was observed by the presence of:

Bilirubin  		≤ 25 mg/dl

Hemoglobin 	≤ 200 mg/dl

Triglycerides  	≤ 800 mg/dl

Ascorbate acid  	≤ 25 mg/dl

For a comprehensive review of interfering substances, refer to the publication by Young.

Quality Controls

It’s necessary, each time the kit is used, to perform quality controls and to check that values obtained are within the acceptance range provided in the insert. Each laboratory should establish its own mean and standard deviation and adopt a quality control program to monitor laboratory testing.

Bibliography

Stein, W. “Creatine kinase (total activity), creatine kinase isoenzymes and variants”. In: Thomas L, ed. Clinical laboratory diagnostics. Frankfurt: TH-Books Verlagsgesellschaft; p. 71-80, (1998).

Recommendations of the German Society for Clinical Chemistry. “Standardization of methods for the estimation of enzyme activities in biological fluids: Standard method for the determination of creatine kinase activity”. J. Clin. Chem. Clin. Biochem., 15, 255-60, (1977).

Moren, L.G., Clin. Chem., 23, 1569 (1977).

Young  D.S., Effects of Drugs on Clinical Laboratory Tests, AACC Press, Washington, DC 5th ed.2000.

Symbols

|  | CE Mark (98/79 CE regulation) |
| --- | --- |
|  | in vitro medical device |
|  | Batch Code |
|  | Use by |
|  | Storage temperature limits |
|  | Read instruction for use |
|  | Gesan Production srl |