Use

Kit for measurement of creatinine in serum, plasma and urine Colorimetric method Jaffé without deproteinization

Summary

Creatinine measurements are used in the diagnosis and treatment of renal diseases and in monitoring renal dialysis.

Principle

Fixed time analysis. Creatinine reacts in alkaline environment with picrate to give a coloured compound whose intensity is proportional to the creatinine concentration in the sample.

Reagents

R1	Litium hydroxide    120.0 mmol/l  Boric acid                    80.0  mmol/l

R2	Picric acid                    67.0 mmol/l

Reagent Preparation

Reagents are liquid and ready to use. About using as monoreagent (“sample starter” procedure) mix the reagents R1 and R2 in equal parts.

Storage and Stability

- 	Store the kit at 15-25°C

- 	After opening, the vials R1, R2  are stable 90 day if recapped immediately and protected from contamination, evaporation, direct light, and stored at the correct temperature.

- 	Working solution stability (R1+ R2): 7 days at 15-25°C.

Precaution In Use

The product is not classified as dangerous (DLg. N. 285 art. 28 l. n. 128/1998). The final concentration of the components, is below the limits imposed by Directives 67/548/EEC and 88/379/EEC and subsequent amendments-to the classification packaging and labeling of dangerous substances.

However the reagent should be handled with caution, according to good laboratory practice.

Waste Management

Please refer to the local legal requirements.

Specimen Collection and Preparation

- 	Serum or plasma. Urine. Diluted 1: 10 (multiply the result by the dilution factor).

-   Do not use hemolyzed samples.

- 	Creatinine in the samples is stable for 7 days when stored at +4 ° C.

Note

- 	The kit, according to this method, must be used in manual procedures. About automatic using follow specific applications.

-	Avoid direct light, contamination and evaporation.

-	The volumes in the procedure can be changed proportionally.

-	In case of complaint or quality control request, refer to the lot number on the package or the lot number on the singles vials.

Procedure

| Wavelength |  |  | λ: 510 (500-550) nm |  |
| --- | --- | --- | --- | --- |
| Working Temperature |  |  | 37°C |  |
| Optical Path |  |  | 1 cm |  |
| Reaction |  | “fixed time” |  |  |
| Monoreagent Procedure “sample starter” |  |  |  |  |
|  | Blank |  | STD | Sample |
| Working Reagent | 1000μl |  | 1000μl | 1000μl |
| Distilled Water | 100 μl |  | – | – |
| Sample | – |  | – | 100 μl |
| Standard | – |  | 100 μl | – |
| Mix and after 30” at 37°C measure the absorbance of sample (E1C) and standard (E1STD) against the reagent blank . After other 2’ make the second reading (E2C), (E2STD). |  |  |  |  |

- Bireagents Procedure “substrate starter”

|  | Blank | STD | Sample |
| --- | --- | --- | --- |
| Reagent R1 | 500μl | 500μl | 500μl |
| Distilled Water | 100 μl | – | – |
| Sample | – | – | 100 μl |
| Standard | – | 100 μl | – |
| Mix, incubate at 37°C for 1’, then add : Reagent R2 500 μl 500 μl 500 μl Mix and after 30’’ at 37°C measure the absorbance of sample (E1C) and standard (E1STD). After other 2’ make the second reading (E2C), (E2STD). |  |  |  |

Calculation

Creatinine [mg/dl]  o [µmol/l]    =

(E2C - E1C)     x      Conc. STD

(E2STD - E1STD)

The reagent performances are related to 37°C,1 cm and 510 nm.

Conversion Factor

Creatinine [mg/dl] x 88.4  = Creatinine [µmol/l]

Reference Values

Serum – plasma

Man 0.9 - 1.3 mg/dl   (80 -115 µmol/l)

Man Urine 800 – 2000 mg/24h (7.1 -17.7 mmol/24h)

Woman  0.6 - 1.1 mg/dl  (53 - 97 µmol/l)

Woman Urine 600 - 1800 mg/24h (7.1 -15.9 mmol/24h)

Reference values are considered indicative since each laboratory should establish reference ranges for its own patient population. The analytical results should be evaluated with other information coming from patient’s clinical history.

ANALYTICAL  PERFORMANCES

The factor and the analytical performance are

reported at 510 nm and 37 ° C

Linearity

The reaction is linear for values ​​between 03 to 25 mg / dl (27-2210 mmol / l). Samples with values ​​exceeding 25 mg / dl must be diluted with saline solution. Multiply, then, the result for diluting factor.

Analytical sensitivity

The test sensitivity in terms of detection limit is  0.3 mg/dl (27 µmol/l).

“Intra-Assay” precision (within-Run)

Determined on 20 samples for each control  (N-H) (Normal-High).

Results:

MEAN [mg/dl]	N =1.15	H = 3.87

S.D.	     N =0.05            H = 0.13

C.V.%	     N = 4.6             H = 3.4

“Inter-Assay” precision (between-Run)

Determined on 20 samples for each control  (H-N).Results:

MEAN [mg/dl]	N = 1.15 	  H = 3.86

S.D.	     N = 0.05	            H = 0.08

C.V.%	     N = 4.19	            H = 2.02

Correlation

A study based comparing this method with a similar method on 20 samples has given a correlating factor  r = 0.98

y=0.98 x + 0.0775

Interferences

No interference was observed by the presence of :

Bilirubin		 ≤ 15mg/dl.

Triglycerides	 ≤ 1500 mg/dl

Hemoglobin		 ≤  150mg/dl

Quality Controls

It’s necessary, each time the kit is used, to make the quality controls and to check that values obtained are within the acceptance range provided in the insert. Each laboratory schould estabilish its own mean and standard deviation and adopt  a quality control program to monitor laboratory testing.

Bibliography

Henry, R.J.: Clinical Chemistry: Principles and Techniques, Harper & Row.N.Y., p. 287, (1964).

Young et al., Clinical Chemistry, 18, (1972).

Newman, D.J., Price, C.P.: Renal function and nitrogen metabolites. In: Burtis CA, Ashwood ER, editors. Tietz Textbook of Clinical Chemistry. 3rd ed. Philadelphia: W.B. Saunders Company; p. 1204, (1999). Kaplan, L.A., Pesce, A..J.: “Clinical Chemistry”, Mosby Ed. (1996).

## Symbols

|  | CE Mark (98/79 CE regulation) |
| --- | --- |
|  | in vitro medical device |
|  | Batch Code |
|  | Use by |
|  | Storage temperature limits |
|  | Read instruction for use |
|  | Gesan Production srl |