Use

Kit for measurement of inorganic phosphorous in serum or plasma. Colorimetric blue of molybdate method.

Summary

Phosphorous measurementes are used in the diagnosis and treatment of disorders characterizedby lowerea levels and elevated levels of inorganic phosphorous.

Principle

End point analysis . Inorganic phosphorous reacts, in acid environment, with ammonium molybdate forming phosphomolybdate complex. The increase in absorbance is proportional to the inorganic phosphorous concentration in the sample.

Reagents

R1	 Hydrochloric acid         1.00 mmol/l

ammonium molybdate   0.70 mmol/l

surface-active agents anionic and

polyanionic.

Reagent Preparation

Reagents are liquid and ready to use.

Storage and stability

-    Store the kit at 15-25°C.

-  After opening, the R1 vial is stable 90 days if recapped immediately and protected from contamination, evaporation, direct light, and stored at the correct temperature.

Precaution in Use

The product is not classified as dangerous (DLg. N. 285 art. 28 l. n. 128/1998). However the reagent should be handled with caution, according to good laboratory practice. Caution: the reagents contain Sodium Azide (0.095%) as preservative. Avoid swallowing and contact with skin, eyes and mucous membranes.

Waste Management

Please refer to the local legal requirements.

Specimen Collection and Preparation

- Serum or plasma. Urine 24h Dilute 1;20

- Do not use samples with haemolysis.

- Do not use sodium fluoride as anticoagulant.

-The phosphorous is stable in the samples up to 7 days at 2-8°C.

Note

- The kit, according to this method, must be used in manual procedures. About automatic using follow specific applications.

- Avoid direct light, contamination and evaporation.

- In case of complaint or quality control request,  refer to the lot number on the package or the lot  number on the single vials.

Procedure

Wavelength : 340 (330-360) nm

Working Temperature 37°C

Optical path 1 cm

Reaction “end point”

-- Monoreagent Procedure “sample starter”

BLANK     STD      SAMPLE

Working Reagent	  1000 µl	    1000 µl 	   1000 µl

Distilled Water	10 µl	--	--

Sample                      	--	            --          10 µl

Standard                        	--	     10 µl	           --

Mix, then incubate for 5’ at 37°C (15-25°C). Measure the absorbance of the sample (EC) and standard (ESTD) against the reagent blank.

Calculation

Phosphorous	 [mg/dl] o [mmol/l] =EC/ESTD  x Conc. STD

The reagent performances are related to 37°C, 1 cm and   340 nm.

Conversion Factor

Phosphorous [mg/dl]   x  0.32  =  Phosphorous [mmol/l]

Reference Values

Serum – plasma - Urine

Adults 2.5 -- 5.5 mg/dl (0.8 – 1.76 mmol/I). Children 4.0 – 7.0 mg/dl (1.28 – 2.24 mmol/I). Urine 300 -1000 mg/24h (96 – 320 mmol/24h)

Reference values are considered indicative since each laboratory should establish reference ranges for its own patient population. The analytical results should be evaluated  with other information coming from patient’s clinical history.

ANALYTICAL PERFORMANCES

Linearity

Reaction is linear up to a concentration of 20 mg/dl Samples with values exceeding 20 mg/dl  must be diluted with saline solution. Then, multiply, the result for diluting factor.

“Intra-Assay” precision (within-Run)

Determined on  20 samples for each control (N-H) (Normal- High). Results:

| MEAN (mg/dl) | N = 3.85 | H = 6.86 |
| --- | --- | --- |
| S.D. | N = 0.06 | H = 0.19 |
| C.V.% | N = 1.62 | H = 2.71 |

“Inter-Assay” precision (between-Run)

Determined on  20 samples for each control (N-H) (Normal- High). Results:

| MEAN (mg/dl) | N = 3.95 | H = 6.58 |
| --- | --- | --- |
| S.D. | N = 0.08 | H = 0.15 |
| C.V.% | N = 2.08 | H = 2.28 |

# Analytical sensitivity

The test sensitivity in terms of detection limit is: 0.10 mg/dl.

Correlation

A study based comparing this method with a similar method on 20 samples has given a correlating factor    r = 0.9956

y = 1.0005x + 0.12

Interferences

No interference was observed by the presence of: Bilirubin		≤  25 mg/dl.

Triglycerides	≤  800 mg/dl.

Hemoglobin 	≤  40 mg/dl.

For a comprehensive review of interfering substances, refer to the publication by Young.

Quality controls

It’s necessary, every time the kit is used, to make the quality controls and to check that values obtained are within the acceptance range provided in the insert.Each laboratori should estabilish its own mean and standard deviation and adopt a quality control program to monitor laboratory testing.

## Bibliography

Kaplan, L.A., Pesce, A..J.: “Clinical Chemistry”, Mosby Ed. (1996).

Talke H, Schubert GE: Klin Wchers., (1965), 43, 174

Young D.S., Effects of Drugs on Clinical Laboratory Tests, AACC Press, Washington, DC 5th ed.2000.

## Simbols

|  | CE Mark (requirement of 98/79 regulation) |
| --- | --- |
|  | in vitro medical device |
|  | Batch Code |
|  | Use by |
|  | Storage temperature limits |
|  | Read instruction for use |
|  | Gesan Production srl |